## Slide 1
Preliminary Design Review
October 14, 2014
Cameron Comeau, Daniel Flora, Smith Johnston, Alexander Potter, Maria Rocco, Evan Schomer, Dylan Stewart, Sarah Wilson, Ashley Zerr

## Slide 2: Agenda
Section | Presenter
Overview | Cameron
Requirements | Dylan
Sensor Selection | Smith
Mechanical | Alexander
Thermal | Daniel, Cameron
Electrical | Dylan
Summary | Smith

## Slide 3: Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 4: Project Statement
The intent is to design, build and validate the functionality of a test bed that will characterize a range of commercial off the shelf (COTS) non-contact infrared temperature sensors, by testing their accuracy and precision over variable distance and temperature and on different target surfaces when compared to contact based temperature sensors, in a simulated space environment.
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 5: Motivation
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Reduce the need for the complex wiring systems.
    Single point failures, expensive and time consuming reliability testing.
  Inability of current sensors to record accurate temperature data where direct contact is not an option.
      Mirrors - Contact sensors would mar and reduce the effectiveness of the surface.
MSL Mast [2]
Multiple wires spanning over moveable joints.
Removes need for contact temperature sensing on moving parts
Notes: Complex wiring … currently used for recording temperature data aboard spacecraft

## Slide 6: Design Overview
Sensor Wheel
Target Wheel
8”
Hollow Shaft
Gear System
Stepper Motor
Targets
ADC
Aerotek Pro 115 Linear Stage (provided by Ball)
Radiator Disk
Insulated rails for Thermal Straps
Cold Finger
(provided by Ball)
50”
IR Sensors
Heater
Copper wires
Heater

## Slide 7: Testbed Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
SENSOR WHEEL
TARGET WHEEL
All dimensions are in INCHES
Thermistor
”
”
”
”
”
”
”
”
”
”

## Slide 8: Concept of Operations
Ball’s TVAC (SIDE VIEW)
Thermistors Pressure Gauge
FRONT VIEW
Target Wheel
Sensor Wheel
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Setup
4 hours
Insert STATIS, De-pressurize TVAC
Takedown
1 hour
Remove STATIS, Re-Pressurize TVAC
Adjust Temperature
Data Collection
Rotate through each Target
Data Collection
Data Collection
Adjust Distance
7 hours for all cycles
TEST CYCLE – Conduct Tests According to Data Collection Cycle
Simultaneous testing of all 4 sensors
Pro115 Mechanical-bearing Ball Screw Linear Stage (Provided by Ball)
3x37 or 2x25 pin Feed-through
Sensor Wheel
Finite Distance Steps
Test CPU
LABVIEW VI
Temp, Distance, Motor
Target ID
Sensor ID Controls/Data Collection
Stepper Motor
Target Wheel
Radiator Plate

## Slide 9: Functional Block Diagram

## Slide 10: Functional Block Diagram - Mechanical

## Slide 11: Functional Block Diagram - Thermal

## Slide 12: Functional Block Diagram - Electrical

## Slide 13: Requirements
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 14: Functional Requirements
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
FR 1: STATIS shall vary parameters (temperature, distance, materials) relevant to characterizing the performance of COTS non-contact infrared temperature sensors.
FR 2: STATIS shall collect measurement data (temperature, distance, materials) required to characterize the performance of the COTS IR sensors.
FR 3: STATIS shall be designed for use within Ball Aerospace’s thermal vacuum chamber.

## Slide 15: Critical Project Feasibility Elements
CPFE.1: Sensor Selection
CPFE.2: Mechanical System
CPFE.3: Thermal Modeling
CPFE.4: Electrical and Software Systems
CPFE.5: TVAC Integration
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 16: Sensor Selection
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Notes: Why are we talking about this first

## Slide 17: CPFE.1 Sensor Selection
The commercial off-the-shelf infrared temperature sensors used in STATIS are critical as they drive:
  Structural design dimensions
    Targets are constrained such that the sensor field of view (FOV) lies entirely within the target.
  Budgetary constraints
    TVAC compatible sensors cost ~$500-800
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 18: IR Sensors– Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Radiation Emitted by Targets
Radiation Absorbed by Detector
Data Output as Voltage
Voltage Converted to Temperature Measurement
Microprocessor
Detector
Non-Contact IR Temperature Sensor

## Slide 19: IR Sensors– D:S Ratio and FOV
IR temperature sensors have a field of view that “captures” the IR radiation.
D:S = Distance : Size ratio
Ranges from 2:1 to >300:1
Field of view (FOV) ranges according to D:S ratio
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
[4]
Worst case 1.5” FOV

## Slide 20: Mechanical System
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Aerotek 115 Linear Stage
Sensor Wheel (translating)
Target Disk (rotating)
Radiator Disk (fixed)
Stepper Motor
Gearing
Hollow Shaft (rotating with targets)

## Slide 21: CPFE.2 Mechanical System
Mechanical system is critical as it drives:
  Manufacturability and geometric tolerancing
    Maintain sensor FOV within target patch.
  Budgetary constraints
    Supply sufficient torque to drive target wheel.
    Sufficient TVAC compatible motors cost ~$700-2000+.
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 22: IR Sensors – FOV Analysis
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Shaft Axis

## Slide 23: Geometric Tolerances
IR sensors’ FOV must fall within target surface.
Largest FOV will govern the size of the targets.
1.5 safety factor on FOV.
Nominal Viewing (units of inches)
Δ =0.375”
Target, 2.25 inch diameter
1.5 inch diameter
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
IR sensor FOV
y
x

## Slide 24: Sensor-Target Alignment Errors
Tolerable 0.06 inch error > 0.005 inch machining precision
FEASIBLE
Elliptical FOV growth (x or y)
y
y
y
x
x
x
x direction shift
y direction shift
=IR sensor FOV
In plane rotation
Wheel tilt

## Slide 25: Torque Consideration
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
τ
Axle
Gear System
Stepper Motor
Target Wheel
Assumptions:
Only Resistance caused by contact in axle bearings
Moment of Inertia is of targets, target wheel, and axle
1:1 Gear Ratio
72 degree rotation angle
0.02
FEASIBLE
Notes: Sensors drive this (FOV – size of target – size of wheel – weight of wheel)

## Slide 26: Thermal System
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Cold Finger

## Slide 27: CPFE.3 Thermal Modeling
Thermal modeling is critical as it drives:
  Thermal rejection and heater design
    Sufficient heat rejection and addition to reach target temperatures.
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 28: Thermal System Layout
Heater
Copper radiator plate
Copper conductivewires
Cold Finger
Cold finger
connectors
Heater
Copper conductive
wires
Copper conduction rails (2x)
Thermistors
Thermistor
Key:
Heater
Conductor
Thermistor

## Slide 29
Heat Transfer Overview
Key:
Radiation
Conduction
qin
qin

## Slide 30: Thermal Circuit Overview - Targets
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Rcond
Contact
Cond.
Temps
Radiation
Rcond
Assumptions:
1D heat transfer
Boundary Conditions: Sensor Wheel: 300 K
  Cold Finger: 50 K
  Shroud: 295 K
Key:
Radiation
Conduction
Contact
Sensorwheel
Cold Finger
Shroud

## Slide 31: Thermal Circuit Overview - Targets
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
EQUIVALENT CIRCUIT:
Assumptions:
All targets and front of target wheel are at the same temperature
1D heat transfer
Boundary Conditions: Sensor Wheel: 300 K
  Cold Finger: 50 K
  Shroud: 295 K
Sensorwheel
Cold finger
Shroud
Target

## Slide 32: Sensors to Targets
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Calculated the rate of heat transfer between the sensor wheel (assumed a constant 300K) and the targets
qrad,sensors
Key:
Radiation
Conduction
sensorwheel
cold finger
Shroud

## Slide 33: Conduction through Targets
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Enforced equal rate of heat transfer on the other side, determined equivalent heater temperature
sensorwheel
qrad,sens
q=-qrad,sens
Theater
Key:
Radiation
Conduction
cold finger
Shroud

## Slide 34: Heater to Cold Finger
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Determine the equivalent rate of heat transfer occurring in the opposite “cold” direction, sum the two sides together to determine total power required by the heater
Cold Finger
Key:
Radiation
Conduction
sensorwheel
cold finger
Shroud

## Slide 35: Heat Required for Targets at Steady State
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Max heater power = 16 W
FEASIBLE.

## Slide 36: Thermal Circuit Overview - Sensors
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Targets
Shroud
Cold finger
Key:
Radiation
Conduction

## Slide 37: Heat Required for Sensors at Steady State
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Max heater power = 13 W
FEASIBLE.

## Slide 38: Cooling Rate
Use of Lumped Capacitance Method
System begins at ambient temperature (T=300K)
System must reach starting temperature (T=250K) within 4 hours (240 minutes)
Component | Time to Cooldown (min)
Target Wheel | 128.4
Graphite Target | 2.4
Glass Target | 7.6
Paint Target | 3.04E-6
Chromated Aluminum Target | 3.42E-6
Anodized Aluminum Target | 3.34E-6
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Cool Down= 128.4 min < 240 min
FEASIBLE.

## Slide 39: Heating Rate
Power required based on chosen time to heat up
  15 minutes for 10 K
  7.5 minutes for 10 K
Times are multiples of minimum specified by DR 3.5.1
  t = time required to heat up
  m = mass of component
  c = specific heat
  Tf = final temperature
  Ti = initial temperature
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 40: Heating Power Requirements
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Max heater power = 40 W
FEASIBLE.

## Slide 41: Electrical and Software Systems
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 42: CPFE.4 Electrical and Software Systems
Electrical system is critical as it drives:
  Target Temperature Measurement
    Thermistor Accuracy = 1K
    Thermistor Precision = 0.1K standard deviation over 30 sec.
    Software system is critical as it drives:
  Schedule
    Development time for software interfaces between workstation, machine controller, and microprocessor
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 43: Electrical Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Notes: Microcontroller external to TVAC
ADC internal to TVAC to reduce signal noise
Stage control handled by separate controller

## Slide 44: Contact Temperature Error
Wheatstone bridge to measure thermistor resistance
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
0.8K
0.4K
Uncalibrated
Calibrated
Requires 16-bit ADC for 0.05K resolution

## Slide 45: Data Budget Summary
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
50B/s

## Slide 46: Power Budget Summary
Chart: Total Power Draw
| Total Power Draw (W)
Main Board | 1.4958
Sensors | 0.06
Heaters | 89.9426
Motors | 12
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Includes 20% Power Margin
104 Watts
Notes: Heater color for total power doesn’t match above

## Slide 47: Software Control
LabVIEW GUI
User control of system
  Target temp setpoint
  Distance setpoint
  Target wheel position setpoint
Record data
Provide feedback to user
Automate testing
Embedded Firmware
Realize control loops
Data acquisition
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 48: Software Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Notes: LabVIEW controls setpoints
Microcontroller realizes feedback control & data acquisition
Data fed back to LabVIEW and saved to CSV
Feedback control for temperature, feed forward control for target wheel position and distance

## Slide 49: Control Interface Mockup
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
User Feedback
User Input
Data Visualization
Test Automation
Record Data

## Slide 50: Summary
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 51: CPFE.5 TVAC Integration
Integration is critical as it drives:
  Financial budget
    All materials entering the thermal vacuum chamber must meet Ball Aerospace’s outgassing specifications.
  Test procedure
    Tests must complete in a 12 hour thermal vacuum testing window.
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 52: Financial Case Studies
Margin
Motor
IR Sensors
CASE 1
CASE 2
Notes: BIGGER ERROR BUDGET…

## Slide 53: Time Budget
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Chart: TVAC Chamber Time
Prep Time | 1
Pump Down | 4
Mechanical | 0.17
Heating | 2.75
Testing | 1.375
Pump Up | 1
Break Down | 1
Buffer | 0.705
Assuming:
30 sec per collection
5 sec per rotation
10 sec per translation
15 min per heat-up
For 11 Temps
  For 3 Distances
    For 5 Targets
Pump Down
Margin
Heating
0.2

## Slide 54: Risk Matrix
3 |  | 5,6 | 
7 | 2 | 8 | 1
4 |  |  | 
Probability of Failure
Importance
High
Low
Low
High
Not finding 4 IR sensors within budget
Not finding a TVAC rated motor within budget
Not finding sufficient heaters for thermal testing
Accurately calibrate heater control system outside TVAC
Purchasing Schedule Risk (order/delivery time frames)
Software Development Time
Access to Ball’s Facilities and Hardware
Thermal modeling accuracy
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 55: Critical Path Moving Forward
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
Thermal Modeling
  Higher dimensional thermal studies, advanced model with access to thermal software
  Motor, Sensor, Heater, Electronics Selection
Electronics Circuit
Integration and Interface Control
Manufacturing and Tolerancing Studies
Major Component Selection
Software Development and Framework

## Slide 56: Questions?

## Slide 57
REFERENCES:
[2] Head of Mast on Mars Rover Curiosity. (2012, August 17). Retrieved October 9, 2014, from http://www.nasa.gov/mission_pages/msl/multimedia/pia15106-anno.html
[3]Boano, C. (n.d.). TempLab: A Testbed Infrastructure to Study the Impact of Temperature on Wireless Sensor Networks. Retrieved September 23, 2014, from http://www.carloalbertoboano.com/documents/boano14templab.pdf
[4] Visible Infrared Imaging Radiometer Suite (VIIRS). (2014, September 14). Retrieved September 15, 2014, from http://npp.gsfc.nasa.gov/viirs.html
[5] Brinton, T. Pentagon Acquisition Chief OKs Weather Satellite Plan. (2010, August 24). Retrieved October 12, 2014, from http://www.spacenews.com/article/pentagon-acquisition-chief-oks-weather-satellite-plan
[6] Student-polished space mirrors ready for launch. (2002, June 24). Retrieved October 12, 2014, from http://spaceflightnow.com/news/n0206/24starshine4/
[7] Compact Non-Contact Infrared Temperature Sensor/Transmitter. Retrieved September 3, 2014, from http://www.omega.com/Temperature/pdf/OS137.pdf
[8] Mclennan VSS/VSH Extreme Environment Stepper Motors. Retrieved October 5, 2014.

## Slide 58: Backup Slides

## Slide 59: Background
TempLab[3] : Testbed created to study the impact of temperature differentials on wireless sensor networks.
  Sensed the accuracy and precision of the wireless components.
Notes:     Utilized infrared heating lamps to heat up electrical components.
    Commanded the temperatures of the components.

## Slide 60: Testbed Overview
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
All dimensions are in INCHES

## Slide 61: Requirements

## Slide 62: Functional Requirement 1
STATIS shall vary parameters relevant to characterizing the performance of COTS non-contact infrared temperature sensors.
DR 1.1 The distance between the target surface and non-contact infrared temperature sensor shall be adjustable through user input commands.
DR 1.2 The temperature of the target surface shall be adjustable through user input commands.
DR 1.3 Every target surface temperature will be measured by every non-contact IR temperature sensor during the allotted test time.

## Slide 63: Functional Requirement 2
STATIS shall collect measurement data required to characterize the performance of the COTS IR sensors.
DR 2.1 The temperatures of the target surfaces shall be collected using conventional contact sensors.
DR 2.2 The distance between the non-contact infrared temperature sensors and targets shall be known.
DR 2.3 The sensor under test (SUT) and target under test (TUT) pair (STP) shall be known.
DR 2.4 Upon completion of any test sequence, a file shall be created that contains data requested in DR 2.2 and DR 2.3.
DR 2.5 All data shall be timestamped relative to the start of each test sequence.

## Slide 64: Functional Requirement 3
STATIS shall be designed for use within Ball Aerospace’s thermal vacuum chamber.
DR 3.1 All required electronic connections between STATIS and test computer shall be compatible with the thermal vacuum cable pass-through.
DR 3.2 The IR temperature sensors shall remain within the temperature range of 275 to 325K.
DR 3.3 The electronics package in the test chamber shall remain within the operating temperature range of the most temperature sensitive device.
DR 3.4 All STATIS components entering the thermal vacuum chamber shall operate within the vacuum environment.
DR 3.5 Testing operations performed by STATIS shall be elected and organized so that testing may be completed within 12 hours inside a thermal vacuum chamber.
DR 3.6 The testbed shall fit within the size constraints of the cylindrical thermal vacuum chamber.

## Slide 65: Sensor Selection - Backup

## Slide 66: IR Sensor Selection Parameters
Field of View (FOV)
  D:S – min 2:1
Cost - 30 – 50% of budget
Geometry
  Similar cylindrical housing
  Smaller sensors can have a cylindrical housing manufactured
Electrical Interface
  adaptable electronics system design
Operating Temperature
  275K – 325K
L = 127mm
D = 24mm
[5]

## Slide 67: Infrared Background
IR radiation has a wavelength ranging from 0.7 to 1000μm (long wavelength)
Depending on the material and sensor capability, wavelengths from 0.7 to 14μm have enough energy to be detected by a range of IR temperature sensors
Moran, Michael J., Howard Shapiro, Bruce Munson, and David DeWitt. Introduction to Thermal Systems Engineering: Thermodynamics, Fluid Mechanics, and Heat Transfer. New York: Wiley, 2003. Print.
Moran, Michael J., Howard Shapiro, Bruce Munson, and David DeWitt. Introduction to Thermal Systems Engineering: Thermodynamics, Fluid Mechanics, and Heat Transfer. New York: Wiley, 2003. Print.

## Slide 68: Emissivity Background
Kirchoff’s Law
  ,
The emissivity of a material is different for different wavelengths
http://www.micro-epsilon.com/download/products/cat--thermoMETER-Infrared--en-us.pdf
Moran, Michael J., Howard Shapiro, Bruce Munson, and David DeWitt. Introduction to Thermal Systems Engineering: Thermodynamics, Fluid Mechanics, and Heat Transfer. New York: Wiley, 2003. Print.

## Slide 69: Geometric Tolerancing

## Slide 70: Geometric Tolerances
Each IR sensors’ FOV should fall only on its’ target surface for the best data collection and sensor characterization. This means that the target array and the IR sensor cluster must be properly aligned.
The sensitivities or tolerances are evaluated in a worst case configuration: maximum viewing distance of 3 inches (largest FOV 2:1).
Nominal Viewing (units of inches)
Δ=R-r
Target, 2.25 inch diameter
IR sensor FOV, 1.5 inch diameter

## Slide 71: Translational Alignment
The difference in the target radius and FOV radius (1.125 in and 0.75 in respectively) gives a bound on the translational alignment between the sensor array and target disk:Δ=R-r, Δ=0.375 inches if it occurs in one direction only within the x-y plane.
x direction sensitivity
y direction sensitivity

## Slide 72: Translational Alignment
Another possibility is that a translational positioning error exists in the x and y directions. In this case, their combined effect must not exceed a magnitude of 0.375 inches.
Assuming the errors in x and y are the same:

## Slide 73: Rotational Alignment: Error in IR Sensor Cluster
Changing the angle of the sensor array about the x or y axes, while keeping the target disk fixed, may be approximated by assuming that a sensors’ FOV translates across the target a distance equal to the magnitude of arc length swept out by a misalignment angle theta.
This implies the approximation s=rθ≅Δ
the angle is being swept out at a 3 inch radius, and must not produce an arc length greater than Δ
s=rθ≅Δ, θ≅Δ/r, θ = ±4.5°
Rotation about y
Rotation about x

## Slide 74: Rotational Alignment: In Plane Rotation (about shaft axis)
Using the same approximation s=rθ≅Δ, the arc length traveled by a misalignment error about the shaft axis (by the target disk) may be approximated as the linear displacement of the FOV off target. Doing similar calculations, but where r is the distance from the center of the target disk to the middle of the target (2.2 inches).
s=rθ≅Δ, θ=±9.6°
Rotation about shaft axis

## Slide 75: Rotational Alignment: Error in IR Target Disk
Keeping the IR sensor array fixed and changing the angle of the target disk about the x or y axes, may be approximated by finding the cross section (projection) of a circular FOV viewing a skewed plane.
This assumes a locally cylindrical FOV,
such that the mapped FOV is an ellipse
has centered between its’ foci
(symmetric about target center).

## Slide 76: Rotational Alignment: Error in IR Target Disk
Rearranging and solving for the angle .
It is found that at dt=2.25 inches
and ds=1.5 inches, the angle at which
the semi-major axis of the mapped
ellipse is larger than the target radius
is 48 degrees, pictured Right.

## Slide 77: Rotational Alignment: Gear Slip
The target array will be driven by a step motor through a gear. The gearing interface provides another possible error source. This error would occur in-plane (about z).
Arc length equivalence:

## Slide 78: Additive Error
Accounting for each error, recall the quadrature sum:
All errors produce displacements in the x-y plane, which gives two perpendicular directions:
recall Δ=0.375 inches
ex and ey each have components discussed on the previous slide. Assuming all errors exist and have the same strength, each must be 0.06 inches or less.
This magnitude applied to the most sensitive error source, sensor array tilt, gives ± 1.2°.

## Slide 79: Geometric Tolerances Summary
Most sensitive translational alignment: 0.27 inches max each in the x-y plane.
Most sensitive rotational alignment: rotation of sensor cluster about x or y axes.
Machining precision: around 5/1000ths of an inch.

## Slide 80: Mechanical Design

## Slide 81: Motor Feasibility
Phytron Extreme Environment Stepper Motor
  VSS HV 32
Vacuum Rated to 10^-7 hPa
With in budget
Provides proper torque and error allowances
  Holding Torque 40mNm,
  Design Voltage 42 VDC
  Step accuracy of 5% for 1.8 degrees
32 mm
11 mm

## Slide 82: Thermal Systems - Backup

## Slide 83: Heater Feasibility
Minco Polymide Thermofoil Patch Heaters
Suitable for vacuum environment
With in budget
Temp Range: 73K – 473K
6-10 in

## Slide 84: Thermal Resistance Equations

## Slide 85: View Factor

## Slide 86: Thermal Circuit Overview – Target Emissivity
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
With the resulting heater power, back-solved for the heat transfer rate through each path and determined the equivalent temperature of each separate target
Contact
Cond.
Temps
Radiation
Rcond

## Slide 87: Sensitivity of Varying Emissivity
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal

## Slide 88: Lumped Capacitance Method

## Slide 89: Lumped Capacitance Radiation Coefficient

## Slide 90: Lumped Capacitance Viability

## Slide 91: Lumped Capacitance Equations
  Must be less than 0.1 to use method

## Slide 92: Electrical and Software Backup Slides

## Slide 93: Contact Temperature Error
Measure differential voltage with reference resistance using a Wheatstone bridge
Requires 16-bit ADC for 0.05K resolution
Absolute Temperature Knowledge 1.0K
  0.8 K maximum error uncalibrated
  0.4 K temperature error with calibration
Electrical noise (0.1K STD) includes ADC, thermal noise floor less than 137 uV (0.05K step size)
Voltage Supply
Error
Thermistor Error
Overview
Sensor Selection
Summary
Requirements
Mechanical
Electrical
Thermal
0.8K
0.4K
Uncalibrated
Calibrated

## Slide 94: Thermistor Wheatstone Bridge
Wheatstone Bridge
Measure differential voltage with reference resistance
Meets 1 K absolute temperature knowledge requirement
Requires 16-bit ADC for 0.05K resolution

## Slide 95: Component Selection Options
ADC | Reference V | Bitdepth
LTC2451 | 2-5V | 16
MCP3421 | 2.048V | 18
ADS1113 | 5V | 16
AD7799 | 2-5V | 24
AD7793 | 2-5V | 24
ADS1242 | 2-5V | 24
ADS1255 | 2-5V | 24
Thermistor | 25C Resistance | Accuracy
PR502J2 | 5k | 0.05C
PR103J2 | 10k | 0.05C
PR503J2 | 50k | 0.05C
PR222J2 | 2.252k | 0.05C
PS103J2 | 10k | 0.1C
Microcontroller | COM | Voltage
SAM3X8E | I2C, SPI, SD, USART/UART | 3.3
ATmega2560 | SPI, UART, I2C | 5
ATSAM4S8CA | I2C, SPI, SD, UART/USART | 3.3
Voltage Reference | Output Voltage | Accuracy
LT1027ECS8 | 5V | 0.10%
ADR01AUJZ | 10V | 0.10%
MAX6126AASA50 | 5V | 0.10%
LM4050AEM3 | 10V | 0.10%
Resistor | Resistance | Accuracy
408-1640-1-ND | 33k ohm | 0.01%
408-1657-1-ND | 47k ohm | 0.01%
408-1654-2-ND | 33k ohm | 0.01%

## Slide 96: Thermistor - Current Source
Current Source

## Slide 97: Workstation Software - LabVIEW
Drivers available
  Stage controller (via Ball Aerospace)
  Microcontroller (LabVIEW LINX library and Arduino toolkit)
Minimal additional user interface programming required
Team familiarity