## Slide 1
Critical Design Review
HoMIE
Holographic Microscope Investigating Enceladus
Notes: Bre

## Slide 2: The Holographic Microscope Investigating Enceladus
HoMIE
Notes: Bre

## Slide 3: The Holographic Microscope Investigating Enceladus
HoMIE
Presenters
Nick
Patricio
Bre
Nick
Eric
Ryan
Freddy
Kellen
Notes: Bre

## Slide 4: Agenda
Design Solution
CPE's & Risk
Design
Requirements & Satisfaction
02
03
04
01
Project
Overview
Spring Planning
Questions
06
07
05
Verification & Validation
Notes: Bre

## Slide 5: Project Overview
Presenter: Brianna Gagliardi
Notes: Bre

## Slide 6
Low-cost astrobiology mission to Enceladus (icy moon of Saturn)
Objective:
  Design, build, and test a flight-like Digital Inline Holographic microscope to support the detection and characterization of life from melted ice grain samples collected from the plumes from Enceladus
Flight-like:
  Design needs to be able to pass basic flight testing (vacuum, thermal, g-loading, etc.), but some aspects of the design will prevent it from being flyable
Enceladus
Encelascope + Transfer Stage
*Not to scale
Project Description
Notes: Bre

## Slide 7: Mission CONOPS
Notes: Bre

## Slide 8: Project CONOPS
Notes: Bre

## Slide 9
*Not to scale
Digital Inline Holographic Microscope (DIHM)
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Interference Pattern Hologram
Reconstructed Image
Notes: Bre

## Slide 10: Design Solution
Presenters: Patricio Ramos, Nick from Electronics, Nick from Software, Brianna Gagliardi
Notes: Patricio

## Slide 11: Full Assembly
Usable 1U Volume
Through Hole in Baseplate and Stand 16 mm by 20 mm
Microscope Stand (Ground only)
10cm
10cm
10cm
Baseplate: ¼” thickness, 140g, 6061 Aluminum
Optical Train Housing
45mm x 57 mm, 142g, 6061 Aluminum
Total weight
(without stand and casing): <500g
Fastener points to spacecraft
Notes: Patricio
Our proposed solution is housed inside a ground assembly, used to keep dust off of our sensitive equipment, and provides a nice base to comfortable route cables and tubings through the metal baseplate hole

## Slide 12: Full Assembly
Syringes used for manual sample insertion and extraction
Microfluidic PTFE tubing used to transport fluid to optical train housing
Epoxy used for gluing the microfluidic tubing to the baseplate
Bubble trap used for preventing air build-up inside tubing
Resistance wire wraps around PTFE tubing for sample heating
Notes: Patricio

## Slide 13: Full Assembly
Image Sensor
Glass Fiber
Borosilicate Sample Tubing
Optical Train Housing
10cm
10cm
10cm
Optical Train Housing Cross Section
Laser Diode Source
Notes: Patricio: Our proposed solution is housed inside a ground assembly, used to keep dust off of our sensitive equipment, and provides a nice base to comfortable route cables and tubings through the metal baseplate hole

## Slide 14: Optical Train
Top-Down View
Image Sensor
Borosilicate Sample Housing
Fastener points to baseplate
57 mm
Lens Stack
Notes: Patricio

## Slide 15: Optical Train
Image Sensor
Borosilicate Sample Housing
Laser Fiber Input
Focusing Lens
Cross-Sectional View
57 mm
Collimating Lens
Aperture
Notes: Patricio

## Slide 16: Software Flow Chart
Notes: Ryan

## Slide 17
Movement Detection Algorithm Flow Chart
Notes: Kellen

## Slide 18: Electronics Functional Block Diagram
Notes: Nick from electronics

## Slide 19: Functional Block Diagram
Notes: Bre

## Slide 20: Risk Reduction
Presenters: Brianna Gagliardi
Notes: Bre

## Slide 21: Critical Project Elements
CPE Number | Name | Challenge
CPE1 | Optical Train | The optical train design must minimize wasted light to meet the resolution requirement.
CPE2 | Sample Heating | The sample must be maintained in a liquid state.
CPE3 | S.W.A.P. | The microscope is limited by flight-like size, weight, and power criteria.
CPE4 | Environment | The microscope must survive flight-like launch and operational conditions.
CPE5 | Hologram Reconstruction Algorithm | The interference patterns captured by the image sensor carries minimal relevant information prior to its reconstruction.
CPE6 | Movement Detection Algorithm | Movement Detection Algorithm must be employed to autonomously search for non-Brownian motion.
CPE7 | Downlink | Meaningful data transmission is limited by the number of bits per second.
Notes: Bre:
For our project, our team worked together to find the six most important elements for delivering a successful product to our customer and their corresponding functional requirements. The ones we decided on were:
Microscope Functionality
  The basis of our whole project, and arguably the most important aspect of the microscope. In order for the image reconstruction and detection algorithms to work, the microscope must be functional to the point where it can take accurate hologram data from the samples.
Microscope Core Interfaces
  This entails the ability for the microscope to handle the interface with the sample collector mechanism at the necessary flow rate for timely material processing. Every bit of water that we are able to process increases the likelihood of detecting Brownian motion and eliminates wasted energy and material.
Spacecraft Interfaces
  This one is pretty self explanatory. The microscope needs to fit within the satellite body and send stored data from the microscope bulk storage to the main processing unit on the bus.
Reconstruction Algorithm
  The reconstruction algorithm is what brings the holographic microscope its special capabilities. This will allow for 3-Dimensional imaging which gives us the ability to more accurately distinguish items within the fluid volume and determine if they are of any interest to us
Detection Algorithm
  Detection is the next step after reconstruction where the computer program will be able to analyze a series of reconstructed images and automatically determine if there are any regions of interest in the sample volume. This reduces the need for high bandwidth downlinks to Earth.
Power
  Finally, the spacecraft only has a limited amount of power that it can generate and distribute through each subsystem. Maintaining a proper power budget not only ensures that the microscope assembly will stay powered, but also protects other essential parts of the spacecraft from losing power in space

## Slide 22: Critical Project Elements
CPE Number | Name | Challenge
CPE1 | Optical Train | The optical train design must minimize wasted light to meet the resolution requirement.
CPE2 | Sample Heating | The sample must be maintained in a liquid state.
CPE3 | S.W.A.P. | The microscope is limited by flight-like size, weight, and power criteria.
CPE4 | Environment | The microscope must survive flight-like launch and operational conditions.
CPE5 | Hologram Reconstruction Algorithm | The interference patterns captured by the image sensor carries minimal relevant information prior to its reconstruction.
CPE6 | Movement Detection Algorithm | Movement Detection Algorithm must be employed to autonomously search for non-Brownian motion.
CPE7 | Downlink | Meaningful data transmission is limited by the number of bits per second.
Notes: Bre:
For our project, our team worked together to find the six most important elements for delivering a successful product to our customer and their corresponding functional requirements. The ones we decided on were:
Microscope Functionality
  The basis of our whole project, and arguably the most important aspect of the microscope. In order for the image reconstruction and detection algorithms to work, the microscope must be functional to the point where it can take accurate hologram data from the samples.
Microscope Core Interfaces
  This entails the ability for the microscope to handle the interface with the sample collector mechanism at the necessary flow rate for timely material processing. Every bit of water that we are able to process increases the likelihood of detecting Brownian motion and eliminates wasted energy and material.
Spacecraft Interfaces
  This one is pretty self explanatory. The microscope needs to fit within the satellite body and send stored data from the microscope bulk storage to the main processing unit on the bus.
Reconstruction Algorithm
  The reconstruction algorithm is what brings the holographic microscope its special capabilities. This will allow for 3-Dimensional imaging which gives us the ability to more accurately distinguish items within the fluid volume and determine if they are of any interest to us
Detection Algorithm
  Detection is the next step after reconstruction where the computer program will be able to analyze a series of reconstructed images and automatically determine if there are any regions of interest in the sample volume. This reduces the need for high bandwidth downlinks to Earth.
Power
  Finally, the spacecraft only has a limited amount of power that it can generate and distribute through each subsystem. Maintaining a proper power budget not only ensures that the microscope assembly will stay powered, but also protects other essential parts of the spacecraft from losing power in space

## Slide 23: Risk Measurement
Impact/Consequence | 
Level | Description
Negligible | No Impact on Microscope Functionality
Minor | Little Impact on Microscope Functionality
Moderate | Negative Impact on Microscope Functionality
Significant | Negative Impact on Microscope Functionality with Isolated Failure
Severe | Negative Impact on Microscope Functionality with Cascading Failures
Probability | 
Level | Description
Very Unlikely | No Probability of Occurring During Microscope Operation
Unlikely | Low Probability of Occurring During Microscope Operation
Possible | Medium Probability of Occurring During Microscope Operation
Likely | High Probability of Occurring During Microscope Operation
Very Likely | Highest Probability of Occurring During Microscope Operation
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Bre:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 24: Critical Project Risk 1
Risk Description | Detection Algorithm Fails to Detect More than 50% of Motion
Mitigation Strategy | Test and improve movement detection algorithm (detailed previously)
Related Subsystem | Software
CPE Effected | CPE5: Movement Detection Algorithm
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  | 1 | 
Likely |  |  |  |  | 
Possible |  |  |  | 1 | 
Unlikely |  |  |  |  | 
Very Unlikely |  |  |  |  | 
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Kellen:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 25: Critical Project Risk 2
Risk Description | Unknown Composition of Sample from Enceladus Requiring More Heat for The Sample
Mitigation Strategy | 1.647 Watts Available to Increase Heater Output to Ensure Sample Remains Liquid
Related Subsystem | Mechanical, Optical
CPE Effected | CPE1: Optics Train, CPE2: Environment
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  |  | 
Likely |  |  |  | 2 | 
Possible |  |  | 2 |  | 
Unlikely |  |  |  |  | 
Very Unlikely |  |  |  |  | 
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Bre:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 26: Critical Project Risk 3
Risk Description | Thermal Damage to Sample and the Electronics Resulting in Component Death
Mitigation Strategy | Thermal Network Modeling to Prevent Heat Buildup
Related Subsystem | Mechanical, Electrical, Optical, Software
CPE Effected | All
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  |  | 
Likely |  |  |  |  | 
Possible |  |  |  |  | 3
Unlikely |  |  |  |  | 
Very Unlikely |  |  |  | 3 | 
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Bre:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 27: Critical Project Risk 4
Risk Description | Sample Housing Becomes Misaligned During Launch Vibrational Environment by 0.5 mm
Mitigation Strategy | 1. Vibrations Simulations & Redesign 2. Fixed Sample Housing Mount in Z-Direction
Related Subsystem | Mechanical, Optical
CPE Effected | All
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  |  | 
Likely |  |  |  |  | 
Possible |  |  |  |  | 
Unlikely |  |  |  |  | 4
Very Unlikely |  |  |  |  | 4
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Eric:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 28: Critical Project Risk 5
Risk Description | Edge Effects of Focusing Lens are Larger Than Anticipated
Mitigation Strategy | Inclusion of 2 mm Circular Aperture Prior to Focusing Lens to Block Edge Effects
Related Subsystem | Optical
CPE Effected | CPE 1: Optical Train
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  |  | 
Likely |  |  |  |  | 
Possible |  |  | 5 |  | 
Unlikely |  | 5 |  |  | 
Very Unlikely |  |  |  |  | 
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Eric:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 29: Critical Project Risk 6
Risk Description | Glass Sample Volume Shatters Under Vibrational Environment
Mitigation Strategy | Mechanical Vibrations Simulations & Redesign
Related Subsystem | Mechanical, Optical
CPE Effected | CPE1: Optical Train, CPE4: Reconstruction Algorithm, CPE5: Movement Detection Algorithm
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  |  | 
Likely |  |  |  |  | 
Possible |  |  |  |  | 6
Unlikely |  |  |  |  | 6
Very Unlikely |  |  |  |  | 
Probability
Impact/Consequence
Acceptable | Watch | Unacceptable
Risk Level:
Notes: Patricio:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 30: Risk Matrix Summary
 | Negligible | Minor | Moderate | Significant | Severe
Very Likely |  |  |  | 1 | 
Likely |  |  |  | 2 | 
Possible |  |  | 2,5 | 1 | 3,6
Unlikely |  | 5 |  |  | 4,6
Very Unlikely |  |  |  | 3 | 4
Acceptable | Watch | Unacceptable
Risk Level:
Impact/Consequence
Probability
Risk Number | Risk Description
1 | Detection Algorithm Fails to Detect Motion
2 | Unknown Composition of Sample
3 | Thermal Damage to Sample and the Electronic Components
4 | Sample Housing Becomes Misaligned During Launch Vibrational Environment
5 | Edge Effects of Focus Lenses are Larger Than Anticipated
6 | Glass Sample Volume Shatters Under Vibrational Environment
Notes: Bre:
Here is our chart for the determined high-risk items on our project. We settled seven of what we deemed the most critical parts.
ACCEPTABLE
False positive from detection algorithm
  Likely, but not bad. Can be addressed with human review back on Earth
FOD from spacecraft bus
  Unlikely given the care taken to prevent FOD in satellite manufacturing/assembly.
  Could interfere with our optics system
FOD in sample collection
  Very unlikely, but if something that wasn’t ice or on a microscopic scale were to be introduced into the collector, it could potentially clog or damage the microfluidic tubing and sample imaging
WATCH
Detection algorithm fails to detect motion
  Very likely due to our inexperience as a team with motion detection code on spacecraft processing platforms
  Minor impact, we can update the detection code remotely or downlink all of our data
Misaligned optical assembly
  Unlikely with proper design and testing
  Severe impact and could ruin the mission for the microscope
Perforation in tubing
  Very unlikely to happen at such a scale
  Severe impact with the possibility of fluid transfer being immobilized
UNACCEPTABLE
Radiation damage to electronics
  Not a worry for our senior projects
  Likely to happen with our electronics choice and radiation environment near Saturn
  Significant impact on the reliability and capability of our reconstruction and detection systems

## Slide 31: Design Requirements & Satisfaction
Presenter: Brianna Gagliardi, Eric Dean, Alfredo Bautista, Ryan Jones

## Slide 32
Number | Name | Driving Requirement
CPE1 | Optical Train | The microscope shall provide a reconstruction resolution of < 0.8μm in the x and y directions, and < 2μm in the z-direction. Note that reconstruction resolution may be a configurable parameter.
CPE2 | Sample Heating | The microscope shall have a liquid water interface for receiving and expelling samples under expected operating conditions.
CPE3 | S.W.A.P. | The microscope payload shall fit within a 1U SWAP (size, weight, and power) allocation.
CPE4 | Environment | The microscope shall survive launch and operational thermal, acceleration, depressurization, and vibrational simulations, requirements
CPE5 | Hologram Reconstruction Algorithm | The microscope’s processor and reconstruction algorithm shall be fast enough to keep up with the rate of material collection.
CPE6 | Movement Detection Algorithm | The software shall provide a method to characterize detected regions and output the characterization as a compact description.
CPE7 | Downlink | The microscope shall support meeting the mission requirements under a 1 kbps downlink constraint.
Driving Requirements
Notes: Bre

## Slide 33
Number | Name | Driving Requirement
CPE1 | Optical Train | The microscope shall provide a reconstruction resolution of < 0.8μm in the x and y directions, and < 2μm in the z-direction. Note that reconstruction resolution may be a configurable parameter.
CPE2 | Sample Heating | The microscope shall have a liquid water interface for receiving and expelling samples under expected operating conditions.
CPE3 | S.W.A.P. | The microscope payload shall fit within a 1U SWAP (size, weight, and power) allocation.
CPE4 | Environment | The microscope shall survive launch and operational thermal, acceleration, depressurization, and vibrational simulations, requirements
CPE5 | Hologram Reconstruction Algorithm | The microscope’s processor and reconstruction algorithm shall be fast enough to keep up with the rate of material collection.
CPE6 | Movement Detection Algorithm | The software shall provide a method to characterize detected regions and output the characterization as a compact description.
CPE7 | Downlink | The microscope shall support meeting the mission requirements under a 1 kbps downlink constraint.
Driving Requirements
Notes: Bre

## Slide 34: Optical Train
R-DIHM-PERF-3: | The microscope shall provide a reconstruction resolution of < 0.8μm in the x and y directions, and < 2μm in the z-direction.
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
R-DIHM-PERF-3.1.2: | The laser shall have a minimum output power of 2mW and an input voltage of 5V.
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Minimum Pixel Resolution: 20482
Minimum Light Intensity: 7.663e-05 [mW/mm2]
Minimum Output Power: 2mW
Input Voltage: 5V
*Not to scale
Notes: Fred

## Slide 35: Optical Train
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Collimator lens effective focal length (mm)
Focus lens effective focal length (mm)
Focus point to detector
distance(mm)
Sample to detector distance (mm)
2x Divergence Angle
(deg)
Collimator/Focus Lens Distance (mm)
*Not to scale
f
s
Notes: fred

## Slide 36: Optical Train
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Collimator lens effective focal length (mm)
Focus lens effective focal length (mm)
Focus point to detector
distance(mm)
Sample to detector distance (mm)
2x Divergence Angle
(deg)
Collimator/Focus Lens Distance (mm)
*Not to scale
f
s
Notes: fred

## Slide 37: Optical Train: Image Sensor Pixel Resolution
Imaging volume
Image sensor
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
Sample to Detector Distance [mm]
Focus Point to Detector Distance [mm]
*Not to scale
f
s
Notes: fred

## Slide 38: Optics Train: Resolution Model
Best Resolution occurs at Turning Point
Pixels can't resolve circular Fresnel fringes at detector
array’s periphery
Detector resolves Fresnel Fringes and acts as an incoherent beam
Fresnel Diffraction
Notes: fred

## Slide 39: Optical Train: Image Sensor Pixel Resolution
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
Turning Point Resolution:
= 405nm
f = 21mm
Notes: eric

## Slide 40: Optical Train: Image Sensor Pixel Resolution
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
Coherent
Resolution:
N =2048
p = 5.5 µm
Notes: eric

## Slide 41: Optical Train: Image Sensor Pixel Resolution
Imaging volume
Image sensor
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
*Not to scale
f = 21 mm
Focus point to detector
distance(mm)
S= 18.6 mm
Sample to detector distance (mm)
Notes: eric

## Slide 42: Optical Train: Space-Grade Image Sensor
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
3D-Plus 4 Megapixel CMOS Space Camera 3DCM734-1
11.3 mm
11.3 mm
5.5 µm
*Not to scale
Notes: eric
This sensor was brought to our attention by our customer as a design point. Our customer wants parts that have space grade equivalents, with the important parameters for us to match being resolution and pixel size

## Slide 43: Optical Train: Prototype Image Sensor
R-DIHM-PERF-3.1: | The image sensor shall have a minimum pixel resolution of 20482
Pixelink PL-D734MU-BL
11.3 mm
11.3 mm
5.5µm
5.5 µm
Maximum Frame Rate: 90.3 FPS
Exposure Time: 28 us to 5s
Power Consumption: 2.88 W
Operating Temperature: 0 to 50 C
Price: $1600
*Not to scale
Manufacturer can send it to us without lens mount
Notes: eric

## Slide 44: Sample Heating
R-DIHM-PHYS-7: | The microscope shall have a liquid water interface for receiving and expelling samples under expected operating conditions.
R-DIHM-PHYS-7.5: | The microscope shall not heat the sample by more than +5o Celsius relative to the baseplate temperature, except when the sample heater is activated.
R-DIHM-PHYS-7.6: | The microscope shall provide a method to actively heat the sample by a configurable amount relative to the baseplate temperature, with 1o Celsius resolution and 1o Celsius accuracy.
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0o Celsius and 5o Celsius .
R-DIHM-PHYS-4: | The microscope shall survive launch and operational environmental conditions.
R-DIHM-PHYS-4.6: | The microscope shall survive temperatures from -50C to +100C.
R-DIHM-PHYS-4.7: | The microscope shall operate over temperatures ranging from -15C to +40C. ​
Notes: Bre

## Slide 45: Sample Heating: Thermal Diagram
Spacecraft baseplate*
Payload baseplate
Optical housing
Radiation from laser
Radiation from image sensor
Conduction from heater
*The Spacecraft baseplate temperature can be set to our desired temperature in order to meet our sample heating requirements
Conduction from spacecraft baseplate
Sample and sample housing
Instrument cross-section
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
Notes: Bre

## Slide 46: Sample Heating: Model
Governing Principles | Assumptions | Governing Equations
1st Law of Thermodynamics |  | 
Conservation of Mass |  | 
Conservation of Momentum | V = 0 | 
 | Heat transfer is instantaneous | 
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
Notes: Bre

## Slide 47: Sample Heating: Survival-Max Temperature Case
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
R-DIHM-PHYS-4.6: | The microscope shall survive temperatures from -50C to +100C.
No laser, image sensor radiation
Sample temperature = 2.5 deg C
  Environment hotter than sample
  Heat needs to be conducted away from sample by baseplate
  Sample heater INACTIVE
  Heat radiated out into space
= 1.4574 W
2.4743 C
Notes: Bre

## Slide 48: Sample Heating: Survival-Min Temperature Case
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
R-DIHM-PHYS-4.6: | The microscope shall survive temperatures from -50C to +100C.
No laser, image sensor radiation
Sample temperature = 2.5 deg C
  Environment is colder than sample temperature
  NO baseplate conduction
  Sample heater ACTIVE
  Heat radiated out into space
= 0.3526 W
= -50 + 273.15 K
2 W available for use
Notes: Bre

## Slide 49: Sample Heating: Operational-Max Temperature Case
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
R-DIHM-PHYS-4.7: | The microscope shall operate over temperatures ranging from -15C to +40C. ​
Sample temperature = 2.5 deg C
  Heat needs to be conducted away from sample by baseplate
  Sample heater INACTIVE
  Heat radiated out into space
  Heat absorbed from image sensor and laser
= 0.4081 W
2.4928 C
Notes: Bre

## Slide 50: Sample Heating: Operational-Min Temperature Case
R-DIHM-PHYS-7.6.1: | The sample temperature shall remain between 0 degrees C and 5 degrees C.
R-DIHM-PHYS-4.7: | The microscope shall operate over temperatures ranging from -15C to +40C. ​
Sample temperature = 2.5 deg C
  NO baseplate conduction
  Sample heater ACTIVE
  Heat radiated out into space
  Heat absorbed from image sensor and laser
= 0.1393 W
2 W available for use
Notes: Bre

## Slide 51: Sample Heating: Future Work
Thermal Desktop Modeling
    More accurate
    Will be able to better account for interactions with aluminum optical housing and glass sample housing
    Predicting it to show we need to add insulation to sample and increase heater output
      Insulation option:
        PVC
Insulation
Sample Tubing
Sample
Notes: Bre

## Slide 52: Verification & Validation
Presenters: Patricio Ramos

## Slide 53: Optical Train Verification
Plan | Place size standard microspheres in imaging volume Capture images with image sensor Reconstruct images Verify reconstructed resolution with software
Driving Requirements | R-DIHM-PERF-3: The microscope shall provide reconstruction resolution of < 0.8μm in the x and y directions, and < 2μm in the z-direction.
Required Items | Laser, collimator lens, focus lens, sample housing, image sensor, image processor.
Notes: Patricio

## Slide 54: Optical Train Verification: Test Setup
Collimating Lens
Focusing Lens
Size Standard Microspheres
Image Sensor
Processor
Data Storage
Data Storage Accessed for Reconstructed Resolution Verification
*Not to scale
Laser
Notes: Patricio

## Slide 55: Sample Heating Verification
Plan | Place thermocouple before, on, and after assembled sample volume Allow thermal equilibrium to be reached for minimum operational temperature bounds Verify heater output and temperature of sample volume Allow thermal equilibrium to be reached for maximum operational temperature bounds Verify heater output and temperature of sample volume Repeat 1-6 for survival temperature bounds
Driving Requirements | R-DIHM-PHYS-7.5: The microscope shall not heat the sample by more than +5 degrees Celsius relative to the baseplate temperature, except when the sample heater is activated. R-DIHM-PHYS-7.6: The microscope shall provide a method to actively heat the sample by a configurable amount relative to the baseplate temperature, with 1 degree Celsius resolution and 1 degree Celsius accuracy.
Required Sensors | Thermistors - NTC 10KOHM 3976K Bead Multimeter – PILOT Lab
Notes: Patricio

## Slide 56: Sample Heating Verification
Instrument Cross-Section View
Heater
Thermistors
Imaging volume
Air temperature reading from environmental condition
Sample & baseplate temperature reading from thermistors
Sample heater power output
Expected sample temperature reading at steady state
Compare
Multimeter
Microcontroller
Notes: Patricio

## Slide 57: Movement Detection Verification
Plan | Use 'test_DHM_2D' to generate test data sets with various classifications of movement Use detection algorithm on test data set to classify data sets Compare classification results of detection algorithm with the known manual classifications of generated data
Driving Requirements | R-DIHM-FUNC-1.5: The software shall provide a method to detect features-of-interest in a single frame reconstruction. R-DIHM-FUNC-1.6: The software shall provide a method to detect features-of-interest in a difference frame reconstruction. R-DIHM-FUNC-1.7: The software shall provide a method to detect features-of-interest in a difference stack reconstruction.
Required Items | Human manual verification
Notes: Ryan

## Slide 58: Generated Data
A
B
C
D
Notes: Ryan

## Slide 59: Movement Detection Verification
Generated Data
Reconstructed Data
Reconstruction Algorithm
1. Use function to generate test data that follows a specific type of classification
2. Run the test data through the motion detection algorithm
3. See how the algorithm performs in comparison to what we know the classification to be
Known Classifications
Algorithmically
Determined
Classifications
test_DHM_2D
Movement
Detection
Algorithm
?
Notes: Ryan

## Slide 60: Spring Planning
Presenters: Brianna Gagliardi
Notes: Bre

## Slide 61: Master Schedule
Systems Testing
Analysis
Manufacturing & Assembly
Deliverables
Notes: Bre

## Slide 62: Master Schedule
Systems Testing
Analysis
Manufacturing & Assembly
Deliverables
Critical Path
Finish
Start
Notes: Bre

## Slide 63: Major Testing Dates
Test | Goal Date | Testing Location
Overall Function | 3/14 | AERO
Thermal Operation | 3/17 | BIOSERVe
Thermal Survival | 3/24 | BIOSERVe
Depressurization | 3/31 | AERO
Notes: Bre

## Slide 64: Financial Budget Breakdown
Project Area | Allocated Budget | Requested Budget | Margin
Mechanical | $1000.00 | $832 | $168
Electronics | $2000.00 | $2043 | -$43
Optics | $600.00 | $530 | $70
Testing | $400 | $300 | $100
Total | $4000.00 | $3505 | $495
Notes: Bre

## Slide 65: Questions?
Notes: Bre

## Slide 66: Backup Slides

## Slide 67: General assembly: Manufacturing Methods
Optical components are sourced from 3rd parties
Steel fasteners and hex nuts from McMaster-Carr
6061 Aluminum, CNC’d by 3rd party (Xometry)
½ in 6061 Aluminum plate from McMaster-Carr, machined @ AERO
Baseplate heater (for testing only)

## Slide 68: Fluid Connector
Sample housing
1/16” OD to ¼”-28 adapter
Machined block of Acrylic/Polycarbonate
Epoxy sealant around edges

## Slide 69: Design Requirements & Satisfaction

## Slide 70: Optics Train: Design
Coherent Resolution Approximation:
Corrected Coherent Resolution:
Turning Point Resolution Approximation:
Source Plane Resolution:

## Slide 71: Optical Train: Light Intensity
Focusing Lens
Collimating Lens
Collimated
radius (mm)
Focus lens focal length (mm)
Initial waist 'w0' (mm)
Initial Waist at the
Focus Point (mm):
w0 = λEFLf/(πRc)
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
*Not to scale
Laser
Divergence Angle:
θ = asin(NA)
Collimated Radius:
Rc = EFLc tan(θ)
Collimator focal length (mm)
θ
Divergence Angle (rad)
w0
EFLf
EFLc
Rc

## Slide 72: Optical Train: Light Intensity
Focusing Lens
Collimating Lens
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
*Not to scale
Laser
Transmission Efficiencies:
Collimating Lens: 0.9 ???
Focusing Lens: 0.95
Sample Housing: 0.9225
Power at the detector = P0*0.9*0.95*0.9225 = 2.36 mW
P0 = 3mW

## Slide 73: Optical Train: Light Intensity
Sample Volume
Image Sensor
Irradiance (mW/mm2):
Initial Irradiance (mW/mm2):
I0 = P0 / (πw02)
Beam Waist (mm):
Rayleigh Range (mm):
z
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
Initial Waist 'w0' (mm)
Sample Axial Distance 'zs' (mm)
Detector Axial Distance 'zd' (mm)
*Not to scale

## Slide 74: Optical Train: Light Intensity
Sample Volume
Image Sensor
Detector Axial Distance 'zd' (mm)
Sample Axial Distance 'zs' (mm)
Initial Waist 'w0' (mm)
z
Image Sensor Size: 5.64 x 5.64 mm
Min waist at detector = 5.64 mm
Influenced by:
Collimated radius
Focus lens focal length
Detector axial distance
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
*Not to scale

## Slide 75: Optical Train: Light Intensity
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].

## Slide 76: Optical Train: Light Intensity
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].

## Slide 77: Optical Train: Sensor Saturation
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].

## Slide 78: Optical Train: Sensor Saturation
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].

## Slide 79: Optical Train: Light Intensity
Waist Radius at the Detector 'wd' (mm)
Waist Radius at the Detector 'wd' (mm]
Detector Area 'Ad' (mm2)
Power at Detector:
Pd = Id x Ad (mW)
Power Loss:
PLoss = Idπwd2 - Pd (mW)
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
*Not to scale

## Slide 80: Optical Train: Focus Lens Focal Length
R-DIHM-PERF-3.1.1: | The minimum light intensity at the image sensor shall be 7.663e-05 [mW/mm2].
θ = asin(NA_laser) = asin(0.13)
Radius = f_col*tan(θ)*1000 = 7.5*tan(θ)*1000
NA_min = resolution/(0.61*lambda) = 0.4
ENA = sin(atan(radius/f_focus))
Plots on next page show max f_focus (decision for focus lens)
θ
Radius
Focusing Lens
Collimating Lens
The ENA grows with decreasing focal length & decreasing focusing lens diameter.

## Slide 81: Optical Train: Focus Lens Focal Length

## Slide 82: Optical Train: Laser Trade Studies

## Slide 83: Optical Train: Laser
R-DIHM-PERF-3.1.2: | The laser shall have a minimum output power of 2mW and an input power of 5V.
FP SM Coaxial Laser Diode Blue-Violet Fiber Laser
Wavelength: 405 nm
Power Output: 3 mW
Numerical Aperture: 0.13
Cost: $109

## Slide 84: Optical Train: Collimator Trade Studies

## Slide 85: Optical Train: Collimator
Focal Length: 7.5 mm
Diameter: 4.5 mm
Numerical Aperture: 0.5
Transmission Rate: 0.9
Cost: $360.55
FC/APC Adjustable Aspheric Collimator CFC8A-A

## Slide 86: Optical Train: Focus Lens Trade Studies

## Slide 87: Optical Train: Focus Lens
Focal Length: 4.0 mm
Numerical Aperture: 0.6
Diameter: 4.80 mm
Transmission Rate: 0.95
Cost: $162.65
V-Coated Aspheric Lens 357775-405

## Slide 88: Optical Train: Image Sensor Pixel Resolution
*Not to scale
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Focus lens effective focal length (mm)
Sample to detector distance (mm)
Numerical Aperture
(mrad)
Collimator/Focus Lens Distance (mm)
21 mm
18.6 mm
Focus point/detector
distance
Collimator lens effective focal length
Notes: eric

## Slide 89: Optical Train
Laser
Collimating Lens
Focusing Lens
Imaging Volume
Image Sensor
Collimator lens effective focal length (mm) 7.5 mm
Focus lens effective focal length (mm) - 2mm
Focus point/detector
distance 'f ‘ (mm) 21 mm
Sample/detector distance 's' (mm) 18.5 mm
2x Divergence Angle
Collimator/Focus Lens Distance (mm) - 1mm
NA = 0.5
CA= 4.5mm

## Slide 90: Environment
R-DIHM-PHYS-4: | The microscope shall survive launch and operational environmental conditions.
R-DIHM-PHYS-4.1 | The microscope shall survive typical launch random vibration simulations up to specified values.
R-DIHM-PHYS-4.2 | The microscope shall operate and meet all performance requirements in the presence of vibration simulations from the orbiter’s reaction wheels.
R-DIHM-PHYS-4.3 | The microscope shall survive a simulated Overall Sound Pressure Level (OASPL) up to 135dB.
R-DIHM-PHYS-4.4 | The microscope shall simulated survive shock loads up to 1000g.
R-DIHM-PHYS-4.5 | The microscope shall survive continuous g-forces of up to 8g on any axis.
R-DIHM-PHYS-4.6 | The microscope shall survive temperatures from -50C to +100C.
R-DIHM-PHYS-4.7 | The microscope shall operate over temperatures ranging from -15C to +40C.
R-DIHM-PHYS-4.8 | The microscope shall survive depressurization at a rate of 0.5 psi/s.
R-DIHM-PHYS-4.9 | The microscope should not employ any materials with outgassing properties that could contaminate any optics or instruments onboard the orbiter, including optics within the microscope.

## Slide 91: Environment: Launch Vibrations Simulations
R-DIHM-PHYS-4.1 | The microscope shall survive typical launch random vibration simulations up to specified values.
Frequency (Hz) | PSD (g2/Hz)
20 | 0.013
50 | 0.030
800 | 0.030
2000 | 0.013
Method
Run SOLIDWORKS random vibration simulation with the given PSD curve.
Multiply maximum stresses in key areas by 3.
Evaluate whether the 3sigma stress is less than the yield strength.
X
Y
Z

## Slide 92: Environment: Operational Vibrations Simulations
R-DIHM-PHYS-4.2 | The microscope shall operate and meet all performance requirements in the presence of vibration simulations from the orbiter’s reaction wheels.
X-Axis Excitation
X
Y
Z

## Slide 93: Environment: Operational Vibrations Simulations
R-DIHM-PHYS-4.2 | The microscope shall operate and meet all performance requirements in the presence of vibration simulations from the orbiter’s reaction wheels.
Y-Axis Excitation
X
Y
Z

## Slide 94: Environment: Operational Vibrations Simulations
R-DIHM-PHYS-4.2 | The microscope shall operate and meet all performance requirements in the presence of vibration simulations from the orbiter’s reaction wheels.
Z-Axis Excitation
X
Y
Z

## Slide 95: SWAP
R-DIHM-PHYS-3: | The microscope payload shall fit within a 1U SWAP (size, weight, and power) allocation that provides all interfaces with the spacecraft.
R-DIHM-PHYS-3.1: | The microscope shall fit within a 1U payload volume.
R-DIHM-PHYS-3.2: | The microscope shall weigh less than 2 kg.
R-DIHM-PHYS-3.3: | The microscope shall consume less than 10 watts active, and 0.5 watts inactive.
R-DIHM-PHYS-3.4: | The processor shall consume less than 10 watts active, and .5 watts inactive.
R-DIHM-PHYS-3.5: | The microscope processor and laser shall not be included in the 1U volume constraint.

## Slide 96: SWAP: Active Power
R-DIHM-PHYS-3.3: | The microscope shall consume less than 10 watts active, and 0.5 watts inactive.
R-DIHM-PHYS-3.4: | The processor shall consume less than 10 watts active, and 0.5 watts inactive.
Chosen components represent ~95% of power consumption
  Some small components have not been chosen, but will not be an issue
MICROSCOPE POWER BUDGET | 
Component | Power [W]
Image Sensor | 2.8
Laser | 0.3
Laser Driver | 0.392
Sample Heaters | 2
Heater Controller | 0.0325
Total Power | 5.5245
PROCESSOR POWER BUDGET | 
Component | Power [W]
Processor | 7.5
Wake-Up System | 0.0325
Total Power | 7.5325
Notes: Active power consumption is not a concern for microscope or processor, as can be seen
Some components not chosen or accounted for such as the relays. These will not increase power consumption in any impactful way
Inactive power not a concern for microscope, can be completely shut down
Inactive power for processor is a concern, processor will need to be shut down completely, and a secondary, low power system will need to be on standby to reactivate it when necessary

## Slide 97: SWAP: Inactive Power
R-DIHM-PHYS-3.3: | The microscope shall consume less than 10 watts active, and 0.5 watts inactive.
R-DIHM-PHYS-3.4: | The processor shall consume less than 10 watts active, and 0.5 watts inactive.
Secondary low power circuit (0.5 W allocated) to wake up the processor.
  Low power microcontroller* will be monitoring input signals for a wake-up sequence
  Upon receipt, the microcontroller will activate a relay, powering up the processor
    From this point, the main processor will take over and begin waking up the microscope
*ATTiny85 used as reference
Notes: This is our plan to lower inactive power consumption to below 0.5 W.
Low power microcontroller will be monitoring input signals, and wake up the processor when required
From that point forward, the main processor will take over
Shown in the figure is a mockup of what this may look like. We used an ATTiny85 for a reference pinout. At this point we plan on using the Teensy LC, as we will be purchasing one for a different system, allowing for hardware familiarity.

## Slide 98: Hologram Reconstruction Algorithm
R-DIHM-PERF-4: | The microscope’s processor and reconstruction algorithm shall be fast enough to keep up with the rate of material collection.
R-DIHM-PERF-4.1: | The microscope shall provide an instantaneous imaging volume > 2μL.
R-DIHM-PERF-4.2: | Reconstruction of the full 3D imaging volume from one frame or one difference stack shall take less than 12,000 seconds at the required reconstruction resolution.
R-DIHM-PERF-4.2.1: | The processor speed shall be in excess of 0.983 GFLOPS, to allow for sufficient power to execute commands and run the detection algorithm.
Interference Pattern Hologram
Reconstructed Image
Reconstruction Algorithm

## Slide 99: Hologram Reconstruction Algorithm
R-DIHM-PERF-4.1: | The microscope shall provide an instantaneous imaging volume > 2μL.

## Slide 100: Hologram Reconstruction Algorithm
R-DIHM-PERF-4.2: | Reconstruction of the full 3D imaging volume from one frame or one difference stack shall take less than 12,000 seconds at the required reconstruction resolution.
R-DIHM-PERF-4.2.1: | The processor speed shall be in excess of 0.983 GigaFLOPS, to allow for sufficient power to execute commands and run the detection algorithm.
Method:
  Use MATLAB function to compute number of FLOPs to perform full reconstruction
  Divided by time of reconstruction (t = 12,000s)
Results:
  Number of FLOPs = 1.18E13 FLOPs
Speed = (FLOPs/time) = (1.1796*10^13/12000) = 0.983 * 10 ^9 FLOPS = 0.983 GigaFLOPS
Processor Research
  Weaver Computer Engineering Research Group at the University of Maine
    Compiled computing power in GFLOPS of various popular processors and SBCs
Notes: https://web.eece.maine.edu/~vweaver/group/machines.html

## Slide 101: Hologram Reconstruction Algorithm: Processor Choice
Processor/Controller |  |  |  | Jetson Nano Developer Kit | Raspberry Pi 4B | Raspberry Pi 3B | BeagleBone Black Industrial | DragonBoard
CRITERIA | Relevant Objective | Weight | SCALE |  |  |  |  | 
Max Power Consumption | PHYS-3.4 | 20% | 3=Most Supportive, 0=Least Supportive | 0 | 2 | 2 | 3 | 2
Computational Power | PERF-4.2.1 | 25% |  | 3 | 2 | 1 | 0 | 1
Thermal Operation and Survival | PHYS-4.6, PHYS-4.7 | 15% |  | 2 | 2 | 2 | 3 | 1
RAM |  | 15% |  | 2 | 3 | 1 | 0 | 1
Cost |  | 10% |  | 0 | 1 | 3 | 2 | 2
Lead Time/Availability |  | 15% |  | 1 | 1 | 3 | 2 | 1
WEIGHTED TOTALS in % |  | 100% |  | 50% | 63% | 62% | 52% | 43%
Notes: https://web.eece.maine.edu/~vweaver/group/machines.html

## Slide 102: Hologram Reconstruction Algorithm: Processor Choice
Winner was Raspberry Pi 4B, but due to supply chain issues, Raspberry Pi 3B will be used for initial prototyping
  Team already owns a Raspberry Pi 3B, and Raspberry Pi 4Bs are not being sold at retail price
  Team will purchase a Pi 4B when available
Notes: https://web.eece.maine.edu/~vweaver/group/machines.html
Several options were considered, including the NVIDIA Jetson Nano, the Raspberry Pi 3 and 4, the BeagleBone Black Industrial, and more, but ultimately the Pi 4 won. Its relatively easy to acquire, has a plethora of resources and support available, and exceeds our requirements.
One thing to note, is that we will begin our prototyping using a Pi 3, as the team currently already owns one, and Pi 4s are not easy to get at retail price currently. We will transition to a Pi 4 when we are able to acquire one

## Slide 103: Hologram Reconstruction Algorithm Flow Chart
1. Subtract intensity of reference wave
2. Take fast fourier transform
7. Take result from step 3 and multiply by the reference wave
3. Over sample

## Slide 104: Hologram Reconstruction Algorithm Flow Chart
5. Multiply result of step 4 with the apodizer and shift to frequency domain
6. Multiply template wave with the apodizer and shift to frequency domain
7. Multiply results of step 5 and 6 and shift to time domain to get the final image of specified z-slice

## Slide 105
Reconstructed Z-Slices

## Slide 106: Frame A Timestep 1

## Slide 107: Frame B Timestep 2

## Slide 108: Frame C Timestep 3

## Slide 109: Frame D Timestep 4

## Slide 110: Difference Stack Algorithm
A
B
C
D
Difference Stack = (B-A)+(D-C)

## Slide 111: Reconstruction
Unprocessed Difference Stack
Reconstructed Difference Stack
Reconstruction Algorithm

## Slide 112: Verification and Validation

## Slide 113: Environment: G-Loading Verification
Plan | 15 g/tons Centrifuge, CIEST Lab located on Main Campus Max volume: 45.72cm x 44.45 cm x 58.42 cm Will ensure physical design matches simulated predictions
Driving Requirements | R-DIHM-PHYS-4.6: The microscope shall survive continuous g-forces of up to 8g on any axis. R-DIHM-PHYS-4.6.1: The microscope shall survive 8g's for 24 seconds.
Required Sensors | Accelerometers
CIEST 15 g/ton Centrifuge

## Slide 114: Environment: Thermal Verification
Plan | Utilize BIOSERVe oven and freezers Place assembly in oven/freezer Allow thermal equilibrium to be reached Pull microscope out of freezer Operate with control sample and test sample immediately and verify outputs
Driving Requirements | R-DIHM-PHYS-4.6: The microscope shall survive temperatures from -50C to +100C. R-DIHM-PHYS-4.7: The microscope shall operate over temperatures ranging from -15C to +40C.
Required Sensors | thermocouples

## Slide 115: Spring Planning

## Slide 116: Financial Budget Breakdown
Mechanical | Mechanical
Item | Cost
Thermistor(split with other team) | $200.00
Optical housing | $150
Optical Housing Endplate | $100
1U Baseplate | $11.70
Sample Heater (split with other team) | $10.00
Epoxy | $20
Sample housing | $0
Syringes | $10
Tubing | $60
Bubble Trap | $120
Sample Housing - Tubing Adapter | $100
Teflon Tape | $5.00
Total | $832
Electronics | Mechanical
Item | Cost
Image Sensor | $1,600
Processor | $80
Camera GPIO Connector | $1.25
Laser Driver | $106.28
Laser | $120
PCB, Misc Components | $75
Teensy LC x2 | $35
Relays | 25
Total | $2043

## Slide 117: Financial Budget Breakdown
Optics | Mechanical
Item | Cost
Collimator Lens | $361
Focusing Lens | $104
Aperture | $65
Total | $530
Testing | Mechanical
Item | Cost
Safety Equiptment | $200
Size Standard Microspheres | $100
Total | $300

## Slide 118: Data Downlink
Using the motion detection algorithm, descriptors and regions of interest will be transmitted via text and JSON files
Researchers can choose to have reconstructed images transmitted of these regions of interest, greatly reducing the image size.
Images requested can be compressed by a ratio of 10:1 to further reduce file size while maintaining image integrity.
R-DIHM-FUNC-1 | The microscope shall support meeting the mission requirements under a 1 kbps downlink constraint.